# Review of Dissolved Oxygen Data in BioChem

Prepared by Gordana Lazin, BIO, October 29th, 2014

# Scope

This report summarizes findings of the dissolved oxygen data review using dataset extracted from BioChem.  The dataset reviewed consists of 147 missions that include AZMP, AZOMP, Fixed stations and Ground Fish survey missions from 1990 to 2013.  The problems found in the data are identified and documented, and comparison between oxygen data derived with different methods is performed on the filtered dataset.

# Background

The amount of oxygen dissolved in the seawater can be measured using different protocols and can be expressed in many different units. In BioChem there are five dissolved oxygen data types. Each data type involves different measurement procedure and is originally reported in different units. The data types and units are summarized in Table 1.

| METHOD (Data type) | Original Units | Comment |
| --- | --- | --- |
| 'O2_CTD_Molar' | mmol/m3 | SI unit, not original unit for any of the methods |
| 'O2_CTD_mLL' | ml/l | Precise measurement but not accurate, systematic offsets that varies between missions |
| 'O2_Electrode' | % saturation | The least precise and the least accurate method |
| 'O2_Winkler' | ml/l | Most accurate method but not very precise |
| 'O2_Winkler_Auto' | ml/l | Most accurate method but not very precise |

Table 1: Dissolved oxygen data types in BioChem, along with units and description notes.

Issues with the oxygen data in BioChem are following:

- Mixture of data with uncertain units

It is not clear what is the actual unit for each oxygen record in BioChem. Most of the oxygen data is converted from original units to mmol/m3. However, some of the data was not converted but is still reported to be in mmol/m3 unit.

- Unit Conversion

It is not clear what conversion factors were used for converting data to mmol/m3. According to BioChem audit it appears that several different factors were used to convert from ml/l  to mmol/m3. The conversion from %saturation to mmol/m3 involves conversion equation, temperature (T) and salinity (S); the origin of T and S used for conversion is not clear.

- Data Validity

We found indications that some of the original data was not loaded correctly and that the current data in BioChem contain number of bogus records, such as values of unknown origin, repeated values, records loaded with zeroes removed from the numbers etc.

# Dataset Overview

Spatial distribution of the data in this dataset is shown on Figure 1. Current dissolved oxygen data in BioChem for each data type is shown on Figure 2. Most of the data belong to 'O2_CTD_mLL', 'O2_Electrode' data types. The histograms for each data type are shown on Figure 3. Note that each data type has almost bimodal distribution, indicating the data in different units.

Figure 1: Spatial distribution of the oxygen data extracted from BioChem used in this review.

Units are uncertain [?]

Figure 2: Dissolved oxygen values in the BioChem data subset under review for each data type and year. The original units in which the data was collected for each data type are in the titles. The actual units in which the values are stored in BioChem are not certain.

Figure 3: Histogram of dissolved oxygen values in the BioChem data subset. The data type and the unit in which the original data was collected are shown under each plot. The bimodal distribution of values for some data type indicates that the oxygen data is stored in different units.

# Identifying Data Units

As mentioned previously dissolved oxygen values in BioChem are all reportedly in mmol/m3. However this is not the case for all records as some of the data values, reportedly in mmol/m3, are actually in different units. Fortunately, each unit is associated with certain range of values that do not overlap greatly, so unit can be ‘guessed’ from the value itself. The common ranges of the dissolved oxygen in the ocean in different units is given in Table 2, the ranges used for ‘guessing’ the units of certain data value in the Table 3., and the summary of the records for each data type and unit in Table 4.

| Range Definition | Unit: mmol/m3 | Unit: ml/l |
| --- | --- | --- |
| Commonly observed range | 200-300 | 5-7 |
| Possible Range | 0-600 | 0-13 |
| Hypoxia | <62 | <1.4 |
| Severe hypoxia | <22 | <0.5 |

Table 2: Dissolved oxygen ranges in the ocean expressed in different units.

| Approximate Range of Data Values | Probable unit |
| --- | --- |
| 0-14 | ml/l |
| 50-105 | % |
| 105-400 | mmol/m3 |
| 14-50 | Suspect data values for all the datatypes |
| 14-90 | Suspect data values for all except ‘O2_Electrode’ datatype |

Table 3: Data ranges used for ‘guessing’ the units for the oxygen records in BioChem.

|  | No. Records | No. Missions | Missions w. oxy in ml/l | Missions w. oxy in mmol/m3 | Missions w. oxy in % |
| --- | --- | --- | --- | --- | --- |
| Whole dataset | 80625 | 147 | 18 (~15% of records) | 222 (~83% of records) | 8 (~2% of records) |
| O2_CTD_Molar | 2716 | 10 | 0 | 10 | 0 |
| O2_CTD_mLL* | 48139 | 98 | 9 | 88 | 0 |
| O2_Electrode | 14618 | 120 | 0 | 112 | 8 |
| O2_Winkler | 11307 | 16 | 6 | 10 | 0 |
| O2_Winkler_Auto | 3845 | 5 | 3 | 2 | 0 |

Table 4: Number of missions and oxygen data records that appear to be in the units listed in the column headings. All the dissolved oxygen data in BioChem is reported to be in mmol/m3.

# Conversion between the units

The amount of oxygen dissolved in the seawater can be expressed in many different units (mg/L, mL/L, µmol kg, % saturation. The unit used most frequently is mL/L and the SI unit is mmol/m3. The conversion is as following:

mL/L * 44.66 = mmol/m3
mL/L = mg/L * (1 mL/1.42903 mg)
mg/L = mL/L * 1.42903 mg/mL
mg/L mg/kg
mg/L ppm

Reference:  ODMS-Ocean Data Management System, St. Lawrence Global Observatory

http://slgo.ca/app-sgdo/en/docs_reference/variables-unites.html

## Effect of the conversion factor in BioChem

Values for dissolved oxygen collected by CTD oxygen sensor in BioChem are mostly in mmol/m3 units even though the original unit is ml/l .  The conversions have been applied using following relation:

ml/l  * (conversion factor) = mmol/m3

As BioChem audit revealed, different conversion factors were applied to different missions.  The conversion factors appear to be one of the following:  44.5, 44.6, 44.66, or 44.661.

Since we are not sure which conversion factors were applied to each mission I attempted to estimate the maximum error that would be introduced by different factors for typical dissolved oxygen concentrations in the ocean.  The exercise is summarized in Table 5. Maximum oxygen value in BioChem table was 557 mmol/m3.

The conclusion is that for dissolved oxygen concentration in the ocean the conversion factor difference would introduce error in the range 0.0081 ml/l to approximately maximum of 0.05 ml/l.

These errors are negligible except if oxygen is used as water mass tracer for which precise measurements are required (Kumiko Azetsu-Scott, personal communication).

| If value in BioChem in [mmol/m3] is | Conversion Factor | Original Value in [ml/l] would be |
| --- | --- | --- |
| 600 mmol/m3 | 44.5 | 13.4831 ml/l |
| 600 mmol/m3 | 44.661 | 13.4345 ml/l |
|  |  | Difference: 0.0486 ml/l |
| 100 | 44.5 | 2.2472 |
| 100 | 44.661 | 2.2391 |
|  |  | Difference: 0.0081 ml/l |

Table 5: Error estimates due to the different conversion factors.

## Oxygen Electrode data conversion

Dissolved oxygen ‘O2_Electrode’ data type is originally collected in % saturation. Converted values that are stored in BiolSum files are converted to ml/l and micro mol/m3 using in-situ T and S from CTD and excel macro from Peter Strain (Jeff Spry, personal communication):

“Ml/l and umol/l are derived from the laboratory  % saturation given by the bench oxygen meter. A conversion program is used that was created by Peter Strain here at BIO. A DOS program that no longer works with win7.

I do not know the underlying formulas but I could give you the spreadsheet/program to look at if you wish.”

The reference for conversion method is following:

O2 (uM) = 44.66 * (% saturation) * saturation concentration(T,S)  : saturation concentration in ml/l according to Weiss, Deep-Sea Research, v 17, 721 (1970)

Strictly speaking, when converting from % saturation to ml/l one should use potential temperature, not in-situ temperature (Kumiko Azetsu-Scott, personal communication). The difference between the potential and in-situ temperature can be about 0.3˚C at 3000 m depth which can cause the difference in oxygen of 0.05 ml/l. That difference is small but is important if oxygen is used as tracer.  Kumiko has a macro to compute 100% saturation level from  ml/l  and T and S (ocean.xla).

Her opinion is that oxygen should be stored in the measured (original) units. This is a nuisance for the user but there is no uncertainty introduced by conversion.

# Data Validity

In order to check on problems with data in BioChem number of checks were performed on the extracted dataset.

### Duplicate Sample ID

Sample ID’s were examined for each mission and data type. Number of duplicate COLLECTOR_SAMPLE_IDs for ‘O2_CTD_mLL’ and ‘O2_Winkler’ data types was discovered in the dataset and is summarized in Table 6.  The oxygen data values under repeated SampleID are different.

| Method | No. Samples | No. Unique Samples | No. Duplicate Sample ID |
| --- | --- | --- | --- |
| O2_CTD_Molar | 2716 | 2716 | 0 |
| O2_CTD_mLL | 48139 | 46254 | 1885 |
| O2_Electrode | 14618 | 14618 | 0 |
| O2_Winkler | 11307 | 10781 | 526 |
| O2_Winkler_Auto | 3845 | 3845 | 0 |

Table 6: Inventory of unique and duplicate sample IDs for each method.

### Gear Type Checks

The gear type in the dataset included: 'Bottle’, 'Bucket', 'CTD', 'unknown', and 'Net'. The gear type ‘Net’ seemed to be suspect and was associated with 61 records in 3 missions:

‘18HU02032’ 10 records:  ‘O2_CTD_mLL’ (5 records) ,  ‘O2_Winkler_Auto’ (5 records)

‘18HU08009’  5 records, ,  ‘O2_Winkler_Auto’

‘18MF12001’ 46 records:  ‘O2_CTD_mLL’ (23 records) ,  ‘O2_Winkler’(23 records)

### Oxygen DATA_VALUE =0

24 records in 4 missions:

'18HU13004' – 1 record, record= 0 in BiolSum, O2_winkler

'18HU99005' – 5 records, 'O2_CTD_mLL'

'18HU99054' – 8 records O2_Electrode (messed up cruise, with zeroes missing)

'18NE99025' - 1 record 'O2_CTD_mLL'

'18PZ00002' - 9 records, 'O2_CTD_mLL'

## Missions with low 'O2_CTD_mLL'

Following three missions were identified that have low oxygen values for CTD data type:

18TL05545, 18TL05546 (original data TEL2005546_hydro.xls), and 18TL09892

The geographic coverage for those missions is shown on Figure 4. For the mission 18TL055046 CTD oxy data was compared with the original records and found to be correct (Figure 5). No comparison with original data was done for the other two missions.  Comparison with Electrode data on all three cruises suggested possible issues with calibration of CTD oxygen sensor (Figure 6).

Figure 4: Missions with low oxygen records: 18TL09892 (dots, Gulf of St. Lawrence), 18TL05546 (circles, Scotian Shelf), 18TL05545 (squares, Gulf of Maine)

Figure 5: CTD dissolved oxygen data for mission 18TL05546 on Scotian Shelf. The values are suspiciously low but BioChem records (red symbols) do correspond to the original data (blue symbols).

Dissolved Oxygen [mmol/m3]

Dissolved Oxygen [mmol/m3]

Dissolved Oxygen [mmol/m3]

Figure 6: Electrode and CTD oxygen data for mission 18TL05546 (top), 18TL05045 (middle), and 18HUTL09892 (bottom). For all three missions there is a large difference between CTD and Electrode oxygen data, with Electrode seemingly closer to expected values, indicating possible calibration issues for CTD oxygen sensor.

## Missions with Problems

18HU04016 – Duplicate ‘O2_Winkler’ and ‘O2_CTD_mLL’

Figure 7: Winkler and CTD data for 18HU04016. The data for both data types contain records with duplicate Sample IDs where one set of samples is in ml/l and the other in mmol/m3.

This mission showed mixture of low and high values in both CTD and Winkler data. Further investigation revealed that data for each data type must have been loaded twice but in different units.  There is no indication in BioChem that the units are different, they are all listed as in mmol/m3.

CTD data: 1570 records, 785 unique COLLECTOR_SAMPLE_ID (2x785=1570)

Winkler data: 1032 records, 516 unique COLLECTOR_SAMPLE_ID (2x516=1032)

Two different MISSION_SEQ:

20000000003201: values look like in [mmol/m3]  and have  5 decimal places

20000000003165: values look like in [ml/l]  and have  3 decimal places

Conversion between ml/l and mmol/m3 seem to be done using factor 44.66, although dividing data values in mmol/m3 with data in ml/l gave range of conversion factors showed on the histogram (Figure 8). This could be a consequence of rounding.

Figure 8: Histogram showing the distribution of ratio [O2 in mmol/m3] / [O2 in ml/l] which was used to deduce the conversion factor.

18HU00050 – Extra Samples for ‘O2_Electrode

Extra 10 samples ID, missing 1 sample ID, the rest seems fine

| BioChem COLLECTOR_SAMPLE_ID | BioChem DATA_VALUE | BiolSum |
| --- | --- | --- |
| 213371 | 264 | NaN |
| 213372 | 270 | NaN |
| 213373 | 270 | NaN |
| 213374 | 271 | NaN |
| 213375 | 273 | NaN |
| 213376 | 266 | NaN |
| 213377 | 263 | NaN |
| 213378 | 263 | NaN |
| 213379 | 247 | NaN |
| 213380 | 220 | NaN |
| 232887 | NaN | 275 |

Table 7: Ten extra collector sample IDs and data values for the Electrode datatype for mission 18HU00050. Those 10 sample IDs found in BioChem for that mission were not found in the original BiolSum file.

Figure 9: Dissolved oxygen data for mission 18HU00050 for Electrode data type. Note that there are 10 samples with sample IDs radically different from the rest of the sample IDs collected on that mission. Comparison with original data from BiolSums does not contain data for these 10 sample IDs.

18HU94008 – Duplicate ‘O2_Winkler’ samples

10 duplicate Winkler samples (see file Duplicate_Winkler_HU94008.xls)

18HU99054 – Zeroes stripped, repeating patterns

Zeroes were dropped for all original records that contain zeroes. For example original record of 300.5 became 3.5 or original record of 290.7 became 29.7. In addition strange repeating patterns were found in the BioChem data. The findings are summarized in excel file comparison_99054.xlsx

18HU01022 – Duplicates ‘O2_CTD_mLL’

Mission had 2002 records for CTD oxygen but only 1001 unique sample IDs what indicated that all CTD samples are duplicates. Each set is associated with different MISSION_SEQ (20000000002102 and 20000000003203) and has different data values associated with each sample ID (Figure 9).  Two sets of CTD oxygen values in BioChem do not seem to be related (Figure 10).

In attempt to resolve the puzzle I examined original CTD oxygen data from primary and secondary sensor using file 2001022_qat.xls.  It appears that the oxygen data values from BioChem with sequence ending in 3203 correspond somehow to the values of the secondary sensor multiplied by 44.66. This seems true for most of the samples except a section of 25 consecutive samples where the values did not approximately match (Figure 12). It is not clear where the values from those 25 samples were coming from, although there is some similarity in values observed in original data (Table 8).  BioChem data from mission sequence ending in 2102 seem to correspond even less to the original data from the primary oxygen sensor, as there was no consistent factor relating them (Figure 11). Those data also show repeated values for the same range of samples that looked suspect in the first set. This indicates that there might have been problems with loading process.

Figure 10: CTD oxygen data for mission 18HU01022.  Two sets of data are associated with different mission sequences.

Figure 11: Two sets of CTD oxygen data values in BioChem from 18HU01022 do not seem to be related.

Figure 12: Values of oxygen data from BioChem for mission sequence divided by the CTD sensor original data. The left plot shows ratio of the BioChem oxygen with mission sequence ending in 3203 divided by the original data from secondary sensor. The right plot shows ratio of the BioChem oxygen with mission sequence ending in 2102 divided by the original data from primary sensor. Note that the ratio is not exact (which might be due to rounding) and that there is a set of samples that largely deviate from the rest of the data (left plot);  lot of scattering on the right plot indicated that there is no consistent relationship between those two data sources.

Figure 13: BioChem oxygen values for the mission sequence ending with 3203 and original CTD oxygen data multiplied by 44.66 from the secondary sensor. It seems that the original data correspond somewhat to the data in BioChem except for the set of samples within the black circle.

| COLLECTOR_SAMPLE_ID | BC_DATA_VALUE MISSION_SEQ 2102 | Original data Primary[mmol/m3] |  | BC_DATA_VALUE MISSION_SEQ 3203 | Original data Secondary [mmol/m3] |
| --- | --- | --- | --- | --- | --- |
| 237713 | 352.04564 | 359.4196606 |  | 226.4262 | 226.7473054 |
| 237714 | 259.87082 | 361.4673216 |  | 234.55432 | 233.8156436 |
| 237715 | 301.66994 | 301.2343796 |  | 191.3681 | 189.3057012 |
| 237716 | 301.75022 | 301.1857002 |  | 192.97586 | 188.6407138 |
| 237717 | 301.75022 | 301.314321 |  | 192.88654 | 189.8706502 |
| 237718 | 301.75022 | 301.4889416 |  | 194.53896 | 190.3328812 |
| 237719 | 301.75022 | 301.5407472 |  | 194.44964 | 189.065877 |
| 237720 | 301.75022 | 301.539854 |  | 192.84188 | 186.7288192 |
| 237721 | 301.75022 | 301.4420486 |  | 189.98364 | 187.241516 |
| 237722 | 301.75022 | 301.3036026 |  | 188.59918 | 187.3562922 |
| 237723 | 301.75022 | 301.222768 |  | 189.40306 | 187.4255152 |
| 237724 | 301.75022 | 301.2687678 |  | 191.18946 | 189.4267298 |
| 237725 | 301.75022 | 301.1754284 |  | 189.49238 | 190.7254426 |
| 237726 | 301.75022 | 301.3348646 |  | 190.2516 | 190.986257 |
| 237727 | 301.75022 | 301.4250778 |  | 190.9215 | 190.0028438 |
| 237728 | 301.75022 | 301.4156992 |  | 191.01082 | 189.6178746 |
| 237729 | 301.75022 | 301.4983202 |  | 189.26908 | 191.2872654 |
| 237730 | 301.75022 | 301.2620688 |  | 187.48268 | 189.195391 |
| 237731 | 301.75022 | 301.3616606 |  | 187.21472 | 188.5848888 |
| 237732 | 301.75022 | 301.4210584 |  | 187.43802 | 190.1649596 |
| 237733 | 301.75022 | 301.399175 |  | 187.21472 | 192.6217062 |
| 237734 | 301.75022 | 301.4308836 |  | 189.00112 | 194.4532128 |
| 237735 | 301.75022 | 301.3228064 |  | 189.93898 | 194.3866694 |
| 237736 | 301.75022 | 301.4340098 |  | 189.84966 | 193.0665198 |
| 237737 | 301.75022 | 301.3911362 |  | 188.4652 | 192.7208514 |
| 237738 | 301.75022 | 301.2553698 |  | 189.31374 | 191.7079626 |
| 237739 | 305.28254 | 305.4592156 |  | 182.39144 | 182.4084108 |
| 237740 | 303.91332 | 303.8871836 |  | 180.51572 | 180.5027686 |
| 237741 | 301.70562 | 300.8467308 |  | 177.7468 | 177.7767222 |
| 237742 | 297.17426 | 296.5348078 |  | 175.37982 | 175.3816064 |
| 237743 | 293.57058 | 292.0665748 |  | 172.92352 | 172.8627824 |

Table 8: Two sets of oxygen CTD values from BioChem (BC_DATA_VALUE) for mission 19HU01022 and original data for primary and secondary CTD oxygen sensor multiplied by 44.66. Note the patterns for the 25 samples highlighted in yellow: for the BC data from mission sequence ending in 2102 one value is repeating; for mission sequence ending in 3203 there is no correspondence to the original data for the highlighted samples, as there is at for the rest of the records that are not highlighted. Note that we don’t know exact conversion factor that was applied to the original data and this is the reason why BC_DATA_VALUES are not exactly the same as original data.

## Missions with multiple MISSION_SEQ

Mission sequences were checked for each mission and the missions with multiple mission sequences were examined.

| DESCRIPTOR | NAME | MISSION_SEQ | METHOD | Note |
| --- | --- | --- | --- | --- |
| ‘18HU01022’ | 'HUD2001022' | 20000000002102 20000000003203 | 'O2_CTD_mLL' (1001) 'O2_Winkler_Auto' (500), 'O2_CTD_mLL' (1001) | Duplicate CTD records, different values for each set |
| ‘18HU04016’ | 'HUD2004016' | 20000000003201 20000000003165 | ‘O2_Winkler’ (516), ‘O2_CTD_mLL’(785) ‘O2_Winkler’(516), ‘O2_CTD_mLL’(785) | In mmol/m3 In ml/l (duplicates) |
| ‘18HU94008’ | 'JGOFS1994' | 20000000001381 20000000003199 | ‘O2_Winkler’ 480 records ‘O2_Winkler’ 10 records (duplicates) ‘O2_CTD_mLL’, 775 records (2 suspects: sample 140642, value 63.507; sample 140817, value 61.247) | List of duplicate winkler data in file: 18HUD94008_doubles_winkler.xlsx |
| ‘18MF12001’ | 'MLB2012001' | 20000000000657 20000000000675 | ‘O2_CTD_Molar’, ‘O2_Electrode’ ‘O2_Winkler’ | Seems OK |

Table 9: Missions with multiple MISSION_SEQ. For two missions multiple sequences indicated that the data was loaded twice (‘18HU01022’ and ‘18HU04016’), in one case (‘18HU94008’) 10 extra records were loaded in separate sequence, and for one mission different data types were loaded at two occasions which seems legitimate.

# Missions with low number of oxygen records

Missions with less than 50 records are listed in Table 10 . Some of those missions are for fixed stations where number of records is expected to be low.  The missions for which there should be much more data available is shown in red.

| DESCRIPTOR | RECORDS |  | DESCRIPTOR | RECORDS |  | DESCRIPTOR | RECORDS |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 18VA01669 | 17 |  | 18VA07666 | 21 |  | 18VA09666 | 18 |
| 18MH03014 | 6 |  | 18PZ00075 | 43 |  | 18VA05668 | 16 |
| 18VA03666 | 12 |  | 18VA99666 | 6 |  | 18HU04019 | 4 |
| 18HU04005 | 12 |  | 18VA02669 | 5 |  | 18VA06669 | 1 |
| 18VA09668 | 24 |  | 18MH06059 | 3 |  | 18HU07033 | 6 |
| 18VA06666 | 18 |  | 18VA06668 | 32 |  | 18VA02666 | 32 |
| 18VA09669 | 2 |  | 18VA07669 | 3 |  | 18VA04669 | 4 |
| 18VA05669 | 6 |  | 18VA01668 | 25 |  | 18VA08666 | 12 |
| 18VA13666 | 48 |  | 18HU02054 | 3 |  | 18VA03669 | 12 |
| 18HU92053 | 44 |  | 18HU10049 | 27 |  | 18PZ00001 | 10 |
| 18VA00666 | 37 |  | 18HU03021 | 3 |  | 18HU01064 | 3 |
| 18VA01666 | 27 |  | 18VA99669 | 6 |  | 18VA03668 | 32 |
| 18VA05666 | 12 |  | 18VA04666 | 24 |  | 18HU03072 | 13 |
| 18VA08669 | 2 |  | 18VA04668 | 15 |  | 18VA07668 | 28 |

Table 10: Missions with less than 50 oxygen records. Missions that should have larger number of records are highlighted in red.

# Summary of Missions with Problems

Missions with problematic CTD data

| DESCRIPTOR | Problem | Notes |
| --- | --- | --- |
| '18HU01022' | 1001 unique samples, 2002 records | All samples are duplicate; Different values for each set. |
| '18HU04016' | 785 unique samples, 1570 records | All samples are duplicate: One set in ml/l , one in mmol/m3 |
| '18HU07001' | Samples IDs from 2 to 152 | Data values are not repeating |
| '18HU07045' | Samples IDs 1 to 24 | Repeated value of 87.166 |
| '18HU09048' | Samples IDs 1 to 29 | Multiple repeated values |
| '18HU10006' | Samples IDs 1 to 622 | Multiple repeated values |
| '18HU11004' | Samples IDs 1 to 24 | Repeated value of 302.134 |
| ‘18NE03003’ | Sample ID < 0 | Values are not repeating |
|  | Suspect Data |  |
| ‘18TL05545’ | Low data values | 2 suspects: sample 140642, value 63.507; sample 140817, value 61.247 |
| ‘18TL05546’ | Low data values | 31 records < 50 mmol/m3, no original data |

Missions with problematic Winkler and Electrode data

| DESCRIPTOR | Problem | Notes |
| --- | --- | --- |
| '18HU00050' | 10 extra samples with values | No data in BiolSum for that samples (Electrode) |
| '18HU94008' | 10 duplicate Winkler records | 2 MISSION_SEQ for Winkler, 10 duplicates have suspect values |
| ‘18HU99054’ | Dropped zeroes + repeating patterns | Needs to be reloaded |

Data Folders with Original Data:

R:\Science\BIO Data and Information\AZMP_datafiles\MFD_sum\2013

R:\Science\BIO Data and Information\AZMP_datafiles\MFD_sum\data_manager

R:\Science\BIO Data and Information\ORMS_DataManagmentFolders\000_DataManagement

# Comparison of oxygen data from BioChem

The missions with problems were filtered out from the original dataset extracted of BioChem and the rest of the seemingly correct data was used to assess the difference between different methodologies  ( CTD, Winkler,  and Electrode).

## Winkler and CTD comparison

Comparison between Winkler and CTD data is shown on the following plots (Figure 13and Figure 14) . Only 8 missions in this dataset were found to have coincidental CTD and Winkler data. The comparison shows systematic difference between CTD and Winkler that varies with each mission. This is a known problem that is due to drifting CTD calibration. Figure 15 is showing Winkler vs. CTD scatter plot for those 8 missions.

Figure 14: Difference between dissolved oxygen from CTD sensor and Winkler or Winkler_Autofor for missions shown in the plot titles. Note that the CTD oxygen sensor tends to measure lower than Winkler and the agreement between those two vary from mission to mission.

Figure 15: Difference between dissolved oxygen from CTD sensor and Winkler and Winkler_Auto for missions shown in the plot titles.

Figure 16: Dissolved oxygen data collected by CTD sensor vs. Winkler for eight missions with overlapping data. The black line is 1:1 line. Note that Winkler measurements tend to be higher than CTD.

## Electrode and CTD comparison

Electrode oxygen data that were converted to ml/l were compared to the CTD data expressed in the same units.  The difference between CTD and Electrode on Figure 16 shows that the oxygen measured with the Electrode tends to be higher than CTD oxygen measurements. The relationship varies between the missions which can be attributed to the drift of CTD oxygen sensor calibration and large uncertainty associated with Electrode method. Scatter plot on Figure 17 is showing the same tendency.

Figure 17: Difference between dissolved oxygen measured by CTD sensor and the one measured by electrode in ml/l using data for all missions.

Figure 18: Dissolved oxygen data collected by CTD sensor vs. data collected using Electrode for all missions with overlapping data. The black line is 1:1 line and the red line is the linear regression through the origin. Note that electrode measurements tend to be higher than CTD with the slope of the regression line of 1.153.

## Electrode and Winkler comparison

There were only two missions in the dataset that included simultaneous Winkler and Electrode oxygen data (18HU12042 and 18NE13022).  However in BioChem Electrode data was expressed in % saturation and Winkler data in ml/l so direct comparison without conversion was not possible. Comparison was attempted using original Electrode data from BiolSums but it appears that % saturation was not converted to ml/l in BiolSum either.

# Conclusions

- The issues with oxygen data in BioChem that are identified in this work include:

  - Uncertain data units

  - Duplicate Sample ID

  - Negative Sample IDs

  - Extra records of unknown origin

  - Repeated data values

  - Sections of data with bogus records (numbers with zeroes stripped, repeating patterns etc.)

  - Inconsistent rounding of the data

Most of those issues would be mitigated by reloading the data.

- Dataset under examination suggested that substantial sections of oxygen data were not loaded to the database. That includes Winkler data and missions identified to have low number of records.

- Comparison of data collected with different methodology is suggesting a calibration drift in CTD oxygen sensor and uncertainty associated with Electrode data. The issue of drifting calibration is currently being assessed and CTD oxygen data collected on recent cruises will include calibration drift correction.

- Recommendation from scientist is that oxygen should be stored in the measured (original) units. This is a nuisance for the user but there is no uncertainty introduced by the conversion.  Unit conversion application could be possibly developed separately as a part of the data query applications.

## Supporting Documents

Catalogue of plots of all the oxygen data for each mission: All_criuses_plots_doxy.docx

Excel file with the table containing number of records for each data type and mission: summary_sorted2print.xlsx

Excel files with data and plots can be found in the folder:

R:\Science\BIO Data and Information\ForGordana